## Commands V1.3
Last update date: 1/7/2021
Notes: - All commands must end with a newline termination - "echo" means the return is exactly the same as the sent command (can be used to verify command was received correctly) - The katana MCU code loops approximately 50 times per second. Avoid sending commands more frequently than this - Commands beginning with X are passed through to the internal motor controller, mainly for debugging purposes - Commands beginning with D are for when in dummy mode (ie. the MCU with no hardware attached). This simulates certain hardware actions - Ending a command string with a semicolon (;) will supress a response. This is useful for commands that don't need error checking
Notes: - Stage position is lost when controller is powered off. This requires a zeroing of the stage for the position to be correct - The zeroing has a repeatability of less than 3 microns - Stage position will drift by a few microns if microtome cable is unplugged
Command string | Command meaning | Example |  | Comment / explanation | Usage (normal / advanced / debug)
|  | Sent | Returned
X[…] | Commands starting with X will be passed-through to the internal motor controller | X? | X0?:PMD401 V13 | Motor controller will respond according to the datasheet. 'X' commands are not necessary for normal operation, only for debugging
XY23 | Query motor status (direct pass-through to motor controller, hence the different syntax) | XY23 XY23 | XY23:83,1 XY23:100,0 | Target reached in 83ms Target not reached, running for 100ms (i.e. has the motor finished the last requested movement. 1 if yes, 0 if no)
K? | Read device ID | K? | K?:KatanaV1.3 SN:xxxxxx
KE | Read encoder pos | KE | KE:250000 | encoder at 250um
KOAM | Get measured knife oscillation amplitude (pm) | KOAM | KOAM:350 | measured 350pm amplitude
KOAF{osc drive strength) | Force oscillator drive strength (0-255) | KOAF200 | echo | Force the oscillation drive strength to 200 out of 255. This will override any previous KOA command
KOAR{1/0} | Turn resonance mode on or off | KOAR1 / KOAR0 | echo | Turning on resonance mode will fix the oscillator drive strength at its current value, and gradually adjust the oscillator frequency in the direction of increasing amplitude. The frequency will settle on the local maximum amplitude. This command will override a previous KOA command
KOA{amplitude} | Set desired oscillation amplitude in nm | KOA120 | echo | Set desired amplitude to 120nm. This command will fix the current frequency, and adjust the drive strength until the desired amplitude is reached
KO_SCOPE | Returns a snapshot of the oscillator return signal waveform (as a series of values over time) | KO_SCOPE | Sample Time (us) Amplitude 0 2 1332 1 27 1285 2 49 1237 … 511 11982 1334 | This outputs the last measured waveform (updated every MCU cycle), that the MCU uses to calculate the oscillator amplidude. This consists of 512 values, over approx 12ms, with a randomised delay (7us+/-2us) to prevent aliasing artefacts. Each newline will be output after a delay of 15ms.
KO_SWEEP | Returns a series of 1024 amplitude measurements over the range 0.1kHz-80kHz | KO_SWEEP | Frequency (Hz) Amplitude (pm) 100 3168.00 178 3185.33 … 80000 4012.67 | This can be used to characterise the knife holder, to visualise its resonance / anti-resonance frequencies. An unusual looking curve will indicate something wrong with the knife holder. Each newline will be output after a delay of 15ms
KO_SWEEP{nSteps, driveStrength, startFrequency, endFrequency} | Same as KO_SWEEP, but accepting arguments | KO_SWEEP3,200,40000,50000 | Frequency (Hz) Amplitude (pm) 40000 838.10 45000 3352.38 50000 856.36 | As above
KOT | Returns termperature of oscillator drive electronics | KOT | KOT:27.53 | Temperature = 27.53 deg This part of the electronics can become hot depending on oscillator driving conditions. The MCU will limit drive strength to maintin a safe temperature (see below)
KOTH | Returns the current temperature throttling value | KOTH | KOTH:1.00 | This is a multiplier for the oscillator drive strength. E.g. 0.50 means drive strength is half the specified value. 1.00 means not throttling.
KOTHSTART{temperature} | Manually set the temperature (deg C) where throttling starts | KOTHSTART35 | echo | If temperature goes above 35C, drive strength will start to be reduced By default this is 35C
KOTHEND{temperature} | Manually set the temperature (deg C) where fully throttled (drive strength forced to zero) | KOTHEND45 | echo | If temperature reaches 45C, drive strength will be zero By default, this is 40C
KOTHOVR{1/0} | Override the temperature throttling | KOTHOVR1 | echo | Turns on override. No throttling will happen
KO | Set oscillator frequency manually (Hz) | KO40590 | echo | set oscillator to 40.59kHz
KT{pos,speed} | Move to position at speed {nm, waveform steps per sec} | KT250000,100 | echo | Move stage to 250um, speed 100. (100 is a reasonable default speed)
KTR | Quesries whether target has been reached | KTR | KTR:1 KTR:0 | Target reached Target not yet reached
KKM | Returns current position of knife | KKM | KKM3385.76 KKM:interlock! | Knife is at position 3385.76 microns Knife cannot move due to main cable being unplugged (in which case knife will be at clear position)
KKM{pos} | Move knife to a specific position (microns) | KKM3600 | echo | Move to 3600um (4000 is fully cleared, 0 is fully cut)
KKC | Initiates a cut ramp | KKC | echo | Startes knife moving from current position, to the cutting window end position
KKP | Reads if knife movement is in progress | KKP | KKP:0 KKP:1 | Knife is not moving Knife is moving
KS | Stop motor and exit closed-loop mode | KS | echo | If cutting, knife will freeze in position
KST | Stop everything (including cut ramp if in progress) | KST | echo | If cutting, knife will freeze in position
KCS | Set knife cutting speed
KMS{speed} | Set knife movement speed (um/s) | KMS1000 | echo | set speed to 1.0mm/s
KWS | Set the cutting window start position (um) | KWS4000 | echo
KWE | Set the cutting window end position (um) | KWS0 | echo
KMD | Memory dump
KRT | Read all realtime data (encoder pos, knife pos, osc freq (if in resonance mode), osc measured amplitude) | KRT | KRT:250132,1602,45950,36 | This will arrive as a stream of bytes (for high speed transmission). This is intended for realtime UI update only, because it cannot be error checked
KSYA | Returns a load of internal variables for debugging | KSYA
KU0 | Gets the status string of the motor controller | KU0 | KU0:0000000000001000 | Digits on RHS can be 1 or 0, and represent: [comError, encError, voltageError, cmdError, reset, xLimit, script, index, servomode, targetLimit, targetMode, targetReached, parked, overheat, reverse, running] (This command is not implemented for dummy mode)
KCP | Gets cable connection status | KCP | KCP:1 KCP:0 | Cable is connected Cable appears to be disconnected (Cable connection is determined by looking at the values coming from the optical encoder
KINTOVR{1/0} | Override the cable interlock | KINTOVR1 KINTOVR0 | echo | The cable interlock has been overriden. The cable interlock is active. The interlock prevents the cut and oscillator (high voltages) from working when the cable is unpluged, due to their high voltages
KID | Get connection ID | KID | KID:Tablet connection KID:SEM connection | This corresponds to the USB ports on the controller (these are actually interchangeable, but this requires extra drivers on the SEM PC)
KI | Get scan state | KI KI KI | KI:2 KI:1 KI:0 | Ready to image Imaging Not imaging The microscope PC should repeatedly poll this whilst waiting for cutting to finish. When cutting finishes, the tablet PC will set the state to KI2. The tablet PC will repeatedly poll this whilst waiting for imaging to finish. When imaging starts, the SEM PC needs to set this to KI1, and when imaging finishes, to KI0
KI{imaging state} | Set scan state | KI2 KI1 KI0 | echo | Ready to image imaging not imaging
KFZ | Initiate stage zeroing procedure | KFZ | echo | The motor will drive until the stage hits the lower hard stop, at which point that stage position will be set as 0nm. This can take up to 5s This needs to be done any time the controller is power cycled
KFZS | Query stage zeroing status | KFZS | KFZS0 KFZS1 KFZS2 | Stage has not been zeroed Stage zeroing in progress Stage has been zeroed
KL0 - LED off KL1 - Red KL2 - Green KL3 - Blue KL4 - Yellow KL5 - Pink KL6 - Turquoise KL7 - White | Set controller LED colour | KL0 KL1 KL2 KL3 KL4 KL5 KL6 KL7 | echo
KDUMMY | Gets dummy state | KDUMMY | KDUMMY:0 KDUMMY:1 | Dummy mode is off Dummy mode is on
KDUMMY{1/0} | Sets dummy state | KDUMMY0 KDUMMY1 | echo | Turn off dummy mode Turn on dummy mode
KFPS | Gets MCU loops per second | KFPS | KFPS:64 | MCU is running at 64 loops per second
DCU{1/0} | Simulate unplugging the main microtome cable | DCU1 DCU0 | echo | DCU1 simulates cable unplugged DCU0 simulates cable plugged (default)
DI{imaging state} | Simulate the imaging state having been changed by the SEM (or other connected PC) | DI2 DI1 DI0 | echo | Ready to image imaging not imaging